## SD14
PRODUCT NAME | 4 MOTORS ELECTRICAL EMERGENCY STRETCHER
PRODUCT BRAND | NITROCARE / NTCR SD14
PRODUCT MODEL | NITROCARE ELECTRICAL STRETCHER
PRODUCT PICTURE
1st MOTOR BRAND, MODEL, ORIGIN | DEWERT / MEGAMAT2 / 2000 N / IPX6 / DC or
LINAK / LA27 / 3500 N / IPX4 / DC
2nd MOTOR BRAND, MODEL, ORIGIN | DEWERT / MEGAMAT2 / 4000 N / IPX6 / DC or
LINAK / LA27 / 6000 N / IPX4 / DC
3rd MOTOR BRAND, MODEL, ORIGIN | DEWERT / GAMMA COLUMN / 2000 N / IPX6 / COLUMN or
LINAK / BL1 / 2000 N / IPX4 / COLUMN
4th MOTOR BRAND, MODEL, ORIGIN | DEWERT / GAMMA COLUMN / 2000 N / IPX6 / COLUMN or
LINAK / BL1 / 2000 N / IPX4 / COLUMN
COLUMN SYSTEM | 2 COLUMN, LINAK or DEWERT BRAND
GENERAL SIZES (BUMPER TO BUMPER) | WIDTH : 990 mm, LENGHT : 2140 mm
HEIGHT MIN | 510 mm
HEIGHT MAX | 910 mm
MOVEMENTS BY MOTORS | BACKREST, LEGREST, HEIGHT, TRENDELENBURG, REVERSE TRENDELENBURG MOVEMENTS
CONTROL PANELS | PATIENT REMOTE CONTROLLER
TRENDELENBURG / REVERSE TRENDELENBURG | TRENDELENBURG / REVERSE TRENDELENBURG BY MOTORS , MAX : 14°
SIDE RAILS ( 2 PCS ) STRUCTURE | PRODUCED FROM CHROM / CHROMED,METAL AND ALUMINIUM MATERIAL, LOCKABLE SIDE RAILS. FOR SAFE USE OF THE STRETCHER, THE STRETCHER HAS HORIZONTALLY FOLDING SIDE SAILS PROTECTORS. THE UPPER PART OF THE SIDE RAILS ARE MADE OF ALUMINUM, THE INTERCONNECTION POINTS ( 5 Pcs - 19 mm) ARE MADE OF STAINLESS STEAL MATERIAL AND THE JOINTS ARE MADE OF PLASTIC MATERIAL.
BED SURFACES | BACKREST HPL ( COMPACT ) for X-RAY
METAL STRUCTURE | THICKNESS FOR BOX PROFILE : 1,2 - 3,00 mm, FOR PIPE 1,2 - 2,00 mm.
COATING TYPE | NANO TECHNOLOGY ELECTROSTATIC POWDER COATING AND COATING THICKNESS MIN : 40 - 90 µm.
MANUAL CPR | MANUAL CPR CAN BE CONTROLLED WITH THE HELP OF HANDLES ON THE RIGHT AND LEFT SIDES.
X-RAY | YES. FULL X-RAY SYSTEM FROM HEAD to FOOT
X-RAY CASSETTE HOLDER | YES
CASTORS | DIAMETER 200 mm / TENTE, HAION or RHOMBUS BRAND. 4 PIECES.
5th CASTOR | 125 mm
LOCKING OF CASTORS | CENTRAL LOCKING SYSTEM, 1.STEP (UP) : DIRECTIONAL CASTOR (YELLOW COLOUR) , 2.STEP (MIDDLE) : FREE CASTORS (FOR ALL CASTORS), 3.STEP ( DOWN) : BRAKE OPEN FOR 4 CASTORS. 2 PEDALS + 1 BAR
MATTRESS | 700 x 1980 x 80 mm, 32 DENSITY, WATER RESISTANT, FIRE RETARDANT MATTRESS (UPHOLSTERY). THE MATTRESS ON THE STRETCHER IS EASILY REMOVABLE, ERASABLE AND HAS A THICKNESS of APPROXIMATELY 80 mm
PROTECTION BUMPERS | 4 PCS BUMPERS ON CORNERS, DIAMETER : 125 mm
BATTERY | YES - STANDART
INFUSION POLE | STAINLESS STEEL, WALL THICKNESS 1,5 mm + PIPE OUTSIDE DIAMETER 25 mm, INNER DIAMETER 19 MM, with 2 PLASTIC HOOKS
DRAINAGE HOOKS & URINE BAG HOLDER | 2 PCS
MANUAL KNEE MECHANISM | KNEE MECHANISM with RATCHET MECHANISM
MAX.SAFE CAPACITY | 275 KG
GUARANTEE | 2 YEARS GENERAL + 5 YEARS for MOTORS
MATTRESS SIZE | 75 x 198 x 8 H
CERTIFICATES | CE, TSE, TUV, TUV THURINGEN, IQNET, ISO 9001:2008, TS 12559, TS 12487, TS 13005, TS EN ISO 13485, TS EN IEC 17025, ISO 14001, OHSAS 18001CE, FDA, SFDA etc.
STANDART SPECIFICATION
• BACKREST, LEGREST, HEIGHT, TRENDELENBURG, REVERSE TRENDELENBURG MOVEMENTS BY MOTORS. • 4 PART STRETCHER SURFACE.• LEG ADJUSTMENT with SPRING MECHANISM • INCLINED BACKREST • 2 PIECES METAL LOCKED AND HIDABLE SIDE RAILS • HYGENIC DESIGN HEAD AND FOOT ARMS MADE FROM METAL MATERIAL.• MECHANICAL CPR ( FLAT POSITION OF BED BACKREST FOR EMERGENCY ) • 360º MOVABLE CENTRAL LOCKING AND LOCAL AND IMPORT CASTORS PROVIDING LINEAR DIRECTIONAL MOVEMENT ( 200 mm ) • PLASTIC PROTECTION BUMPERS • X-RAY • 5.CASTOR MOVEMENT EASINESS TO DESIRED DIRECTIONS • REQUIREMENT BASKET • URINE BAG HOLDER • BELT HOOK • ACCESSORIES SLOTS • HEIGHT ADJUSTING IV POLE • ELECTROSTATIC POWDER COATING • 32 DENSITY MATTRESS.